- Does the malnourished child with improper Body water content affect the 50khz trunk impedance value?

Yes, a malnourished child with improper body water content significantly affects the 50kHz trunk impedance value. In bioelectrical impedance analysis (BIA), impedance is the opposition to an electric current; because water and electrolytes are excellent conductors, changes in their volume and distribution directly alter this measurement.

The trunk is a large, relatively short cylinder of the body. Even though it accounts for nearly 50% of total body weight, it contributes only a small fraction (about 10%) of total body impedance. Therefore, even minor changes in trunk fluid can cause noticeable shifts in the 50kHz impedance value.

- Fluid Overload (Oedematous Malnutrition/Kwashiorkor): Children with nutritional oedema have high levels of extracellular water (ECW) in the trunk. Because water conducts electricity easily, trunk impedance values will be significantly lower than normal.

- Dehydration (Severe Wasting/Marasmus): In children with marasmus or severe wasting who are also dehydrated, the loss of total body water (TBW) increases the resistance to current. In these cases, trunk impedance values will be significantly higher.

- Tissue Loss: Malnutrition involves the loss of muscle mass (the primary reservoir for intracellular water). Less muscle means less conductive fluid, further increasing impedance.

2) Any studies supporting the validation criteria used (50 khz) for checking the data validity? If yes, how can it still be considered for a malnourished child (MAM / SAM) as they have a poor hydration status and fat percentage?

The 50 kHz frequency is the "gold standard" for Single-Frequency BIA (SFBIA) because it represents a point where the electrical current can partially penetrate cell membranes.

- The Cole Model: As discussed in the paper you linked (PMC10140124), 50 kHz is often very close to the Characteristic Frequency ($f_c$) in humans. At this frequency, the Reactance ($X_c$) is near its maximum, allowing for the best possible calculation of the Phase Angle (PhA).

- Validation Studies: Most predictive equations for Lean Body Mass (LBM) were developed and validated at 50 kHz because it provides a reliable balance between measuring extracellular fluid and reflecting cellular mass.

Even though hydration is abnormal in MAM/SAM, the validation criterion at 50 kHz is not used to “confirm normal physiology.” It is used to confirm that:

- The measurement was technically sound

- Electrodes were properly placed

- No major movement artifact occurred

- Current transmission occurred correctly

So the frequency remains valid as a technical reliability checkpoint, even if body composition interpretation requires caution in malnourished populations.

In short:
 Hydration affects impedance values
 But 50 kHz remains appropriate for checking measurement validity
 Interpretation of fat mass and lean mass must still consider altered hydration

3) As per our observation, the data that was valid had a phase angle value of zero. How can that be justifiable?

A phase angle of zero at 50 kHz is physiologically unlikely in a living subject.

Phase angle is calculated from:

- Reactance (Xc)

- Resistance (R)

If phase angle is zero, it means:

- Reactance = 0
or

- Reactance was not recorded

In the case of devices like the InBody S10, phase angle will show as zero if:

- The device was not in research mode

- Reactance measurement was not activated

- The parameter was not exported in that testing configuration

- Software settings suppressed that output

4) Why can't we use phase angle and reactance as a valid measure for checking the data?

Reactance alone cannot validate data quality because it is a measure of cell membrane capacity, not an indicator of the device's technical performance. In malnourished children (MAM/SAM), reactance is naturally low due to damaged cell membranes and severe muscle wasting; if you used low reactance as a reason to reject data, you would be accidentally deleting the very children you are trying to study. To validate data, you must look at Resistance and Reactance together (the Impedance Vector) to ensure the electrical signal is consistent, rather than relying on one value that is biologically expected to be abnormal in this population.